

# Control Settings:

* Tidal Volume – User Settable: 300 to 900 ml
* Respiratory Rate – User Settable: 10 to 34 Bpm (I:E Maximum 1:1; Minimum 1:12)
* Peak Inspired Pressure – User Settable: 15 to 50 cm H2O
* PEEP Level – User Settable: Off, 5 to 20 cm H2O
* Spontaneous Trigger Sensitivity – User Settable: -10 to -2 cm H2O
* O2 % - Discrete Settings: 60%, 80% and 100%
* Inspiratory Flow – Fixed flow at 40 Lpm
* On/Standby switch for suction procedures

# Ventilation Mode

* Volume Controlled Ventilation (VCV)
* Assist Control – Spontaneous Effort Triggered VCV breaths
* Inspiratory Pressure Limited at PIP setting with pause
* Pmax safety limit at 60 cm H2O

# Monitoring & Alarms

* LCD Digital Display of Ppeak, Pmin, Vte and Rate values updated each breath
* Mechanical Pressure Gauge for visual pressure representation
* Battery and Alarm Status continually displayed
* Alarms include High Paw, High & Low Vte, Circuit Disconnect, Occlusion, On Battery
* Alarm audio tone and color visual indicators
* Uses delta pressure type airway flow sensor (GE D-Lite)

# Physical Operation

* Size: 15 cm H x 23 cm W x 27 cm D; Weight: 2.3 Kg
* Electrical Power Input: 12 Vdc Input, 2A (external AC/DC power supply)
* Pneumatic Input: O2 supply at 35 to 87 psi (all O2 delivered to patient)
* Battery Backup: Automatic switchover and recharge with >30 minute operational capacity
* Operating Environment: 10 to 40 C; 15 to 95% RH Non-Condensing
* Storage Environment: -20 to 65 C; 15 to 95% RH Non-Condensing
* Vibration: Meets National Safe Transit Association (NSTA); Procedure 1a
* Standards: Meets ISO 80601-2-12 with IEC 60601-1 Critical Care Ventilator Standards